## Slide 1: March Run Study
Engineering Run 15

## Slide 2: Centroid Repeatilibility

## Slide 3
2023/July
MCS realignment caused the PSF change. Although the overall image qulaity is consistent, the difference in PSF shape impact the centroid algorithm.
2023/July

## Slide 4
There are two issues related to the centroid code, kernel size and culling method.

## Slide 5
We fisrtly turn off the multi-processing.
Then, we changed the kernel size.
The centroid repeatibility is improved.

## Slide 6
The convergence is improved after the correction with the same design.

## Slide 7: Convergence Study

## Slide 8
In the eng15 run, we did 164 convergence runs. However, we spend 54 runs on revising cetroid code. Therefore, we have 110 convergence runs.
There are 87 runs with 75% percentile < 0.01mm.

## Slide 10
Example of bad convergence run

## Slide 11: Cobra Geometry

## Slide 12
EL90 on 3/14

## Slide 13
EL90 on 3/16

## Slide 15: EL60-EL90

## Slide 16: EL30-EL90

## Slide 17: EL30-EL60

## Slide 18: Cobra Armlength

## Slide 19: L1 (theta EL60-EL90)

## Slide 20: L2 (EL60-EL90)

## Slide 21: L1 (EL30-EL90)

## Slide 22: L2 (EL30-EL90)

## Slide 23: L1 (EL30-EL60)

## Slide 24: L2 (EL30-EL60)

## Slide 25
The difference between current and measured L1 arm lengths is getting bigger. We need to update the XML.

## Slide 27: Bad Convergence Cobras

## Slide 28
Among 110 convergence runs, only one cobra (cobra index = 570) repeatedly bad. This means the phi arm of this cobra is not opening due hardware issue.

## Slide 29: Bad Cobra with Same Design

## Slide 30
For studying the non-converged cobra, we checked the convergece runs with same design. Commanding cobra to the same design for 3 ~4 times, we can find out the cobras repeatedly bad.
We sent throught all cobra movement plots and identified serveral cases of non-convergence cobra.
  Bouncing: cobra is bounceing around the target. This is the majority of the convergence cobras (56.9%).
  Need Small Move: cobra is close to target. Those cobras just need very small move to get on targets (seeing? 16.7%).
  Sticky: the phi or theta arm is hardly moving. We have identify phi arm of cobra 570 is hardly open acrsso all designs (5.6%).
  Fiber ID confusion: ~(20.8%)

## Slide 31
On March 15, we took several convergence runs with same design. This is for study the repeatibility of the “not done” (distance > 100 um) cobra.

## Slide 33
Repeatedly Bad Cobra of 108070, 108071, 108072, 108073
Repeatedly Bad Cobra of 108065, 108066, 108068

## Slide 34: Bouncing

## Slide 35: Need small move

## Slide 36: Sticky
Note: Cobra 570 is identified as no phi movement in previsous analysis

## Slide 37: Fiber ID Confusion

## Slide 38: Operation Time

## Slide 39
This is the FPS and MCS operation time in the 2023 December engineering run.

## Slide 40
The MCS operation time is incresed because we turn off the threading function in centroid code. We need to test the centroid code more.